Proposal for the Hall-D Beam Line Pair Spectrometer - Goals,  Technical Description, and Specifications

1) Introduction

The photon beam source for Hall-D will be based on coherent bremsstrahlung of the CEBAF 12 GeV beam with a diamond radiator. The photon beam generated by the interaction with the diamond radiator will consist of contributions both of incoherent un-polarized bremsstrahlung photons and partially linearly polarized coherent bremsstrahlung photons. As the angular divergence of the incoherent  bremsstrahlung  is larger than one for the coherent bremsstrahlung, the photon beam  can be enriched by a coherent component if the tight collimation is applied using enough long drift space from  the diamond radiator. It is planned to place the primary collimator of 3.4mm in diameter  for Hall-D 76 m downstream of the diamond radiator which is roughly 22m in front of the GlueX detector and restricts the polar  angles above  ~22.µrad as compared to a characteristic divergence of the incoherent bremsstrahlung, a me/Ee ≈ 42µrad 

. A tagging spectrometer, referred to as the tagger, will be built 2 m downstream of the radiator and will measure the energy of the electrons which radiated the photons, the photon energy will then determined from energy conservation. The tagger will be readout for every GlueX trigger and will be the primary measurement of the photon energy, but the measured photon flux Iγmeas(E) will be convolution of the collimated photon  spectrum Iγcoll(E) with the trigger efficiency ε(E).    Cross section measurements at GlueX will require a knowledge of  the photon    spectrum Iγcoll(E), that .in principle  could be determined by the GlueX detector using the cross sections for some known process, as well as understanding the acceptance and trigger efficiencies for that reaction, to better than 2% in order to determine the photon flux at the 2% level. from the measured trigger rates. In practice, an additional instrument will be needed for a precise measurement of the photon flux incident on the GlueX target.  In addition to the absolute flux measurement, the linear polarization of the photon beam also needs measured. 
 In order to have a precision measurement of the characteristics of the photon beam it is necessary to base the measurement on a well understood process. It is proposed to use pair production which is a well understood QED process as the basis for the luminosity measurement and polarization determination. The pair production cross section between E( of 6 and  12GeV known to an accuracy better than   1%. In this document a pair spectrometer is proposed based on a  single dipole  magnet and a simple detector array.  Such a  magnet design, with simple field topography, is chosen to minimize the  uncertainty of  the electron and positron trajectories calculated . The pair spectrometer is self calibrating to the extent that the absolute energy scale can be determined from the high energy edge of the photon spectrum. The endpoint of the spectrum corresponds to the CEBAF beam energy which is known to better than 1 part in 10-3.  The beam flux Iγcoll(E)  is determined from the coincidence rates in the spectrometer. The excellent energy calibration of the pair spectrometer will allow a cross calibration of the photon energy measurement in the tagger spectrometer by looking at coincidences between the two spectrometers.   This technique can also be used to determine the tagger trigger efficiency. 
Finally a careful analysis of the flux function Iγcoll(E) can determine the polarization of the photon beam.
 In summary the goals for this Pair Spectrometer are as follows.

      a)  Continuously measure the post-collimated beam flux Iγcoll(E) as a function of energy 
      between  6 and 12 GeV.
b)   Determine the photon beam linear polarization.
c)   Provide an independent tagger energy calibration.
d)  Measure the tagger trigger efficiency.
In addition to the above goals the pair spectrometer must be designed so that the spectrometer can be upgraded in the future with silicon pixel detectors to measure the photon beam linear polarization directly. These detectors must be placed upstream and downstream of the pair spectrometer magnet, and have sufficient separation to resolve both the polar and azimuthal angles of the electron and positron momenta.
e) Pair Spectrometer Design

 [image: image6.jpg]7

9

w01
Epeam = Eragger




Figure 1 Layout of the proposed pair spectrometer. On the left  is a thin 10-4 radiation length converter followed by a strong spectrometer magnet. The vacuum chamber extends 1.5 m beyond the end of the magnet. Immediately outside the vacuum chamber is a scintillator hodoscope detector.
The layout of the proposed pair spectrometer is shown in Figure 1.  The 6 – 12 GeV photon beam is sampled with a thin converter shown on the left of Figure 1.  As the pair production cross section is large, a thin converter can be used.   The optimum choice for the converter thickness will depend on the details of the detector design and photon beam flux and will therefore be discussed after the detector is described. The target holder will be positioned close to the entrance of the pair spectrometer magnet and will be equipped with several different thickness radiators. The converter ladder may also be completely removed from the beam. 

An expected beam spot at the PS converter is 3.4mm in diameter and is much larger as compared to the projection of  opening angles of produced  pairs in the PS detection planes having production angle divergence  me/E( = 42 - 85 µrad and the multiple Coulomb scattering angles in order of 55-145 µrad, assuming converter thickness 10-3Xo. The magnet field needs to be strong enough because the larger the Int[B*dL]  of the magnet the more dispersed the electrons and positrons will be on the hodoscope plane and the better the momentum resolution. The pole needs to be also  wide  enough because the width of the pole defines the lowest momentum measureable in the spectrometer. A  survey of available magnets has been performed and the magnet best suited for our applications selected. This magnet has a  pole width of about 600 mm and an effective length of about 1.83 m. The exact numbers can only be determined after plates have been designed to reduce the magnet gap from 150 mm to 25-30mm ( >10σ, where σ -1mm for the beam spot size of 3.4mm). The magnet shall operate at a field where the iron is not saturated to provide  the fixed field shape, attributed to a linear part of  magnetic field /current dependence  . For this study a 1.64 tesla field is assumed. If the magnet is operated at this field, electrons/positrons with a momentum of 3.0 GeV have a horizontal displacement at the end of the magnet of 28 cm, so the pole width of ±30 cm is just sufficient to accept this momentum. The vacuum chamber downstream of the magnet is limited in length to about 1.5 m by space requirements for equipment necessary for the GlueX detector. Given the parameters of the magnet the chamber needs to be 1.6m wide at the downstream end. One necessary condition on the spectrometer design is that 12 GeV electrons or positrons generated in the converter must be able to pass through the exit window at the end of the vacuum chamber and strike a shielding wall. This protects GlueX from unnecessary background.  With this design 12 GeV electrons or positrons will exit the vacuum chamber 18 cm from photon beam axis leaving sufficient space for a flange to couple the photon beamline and the exit window mount. The vacuum chamber is constructed of non-magnetic material (aluminum or stainless steel). In summary, a magnet with an effective length Leff = 1.829 m, an integrated field strength Int[B*dL]  =2.991 T*m, an operating field Bmax = 1.64 tesla, and a pole width Dx = ±30 cm is assumed. The vacuum chamber extends 1.5 m beyond the end of the magnet.
. 
The scintillator hodoscope detectors are shown in the right of the  Figure 1. The Fine Spacing Forward (FSF) hodoscope (positron arm) and Wide Spacing Forward (WSF) hodoscope (electron arm) are used for the energy measurement while the Fine Spacing Backward hodoscope (FSB) and Wide Spacing Backward hodoscopes (WSB) are used for triggering purposes. The FSF hodoscope will measure in fine momentum bins between 3 and 4 GeV/c. The WSF is a sequence of single scintillators at intervals corresponding to 1GeV and covering the range 3.25 to 8.25 GeV. As the spacing between the counters in the WSF hodoscope corresponds to the total range of coverage in the FSF arm, the every momentum bin from 12.25 GeV to 6.25 GeV corresponds to a coincidence between only one pair of scintillators. This spectrometer design provides uniform acceptance over the energy range of 12.4 GeV to 6.4 GeV using a minimum number of counters and readout channels. The disadvantage of the design is that the overall  acceptance for the produced pairs is low, that is  the cost for the fine photon energy resolution. Cross sections, rates, and acceptance will be discussed in the performance section

[image: image2.jpg]momentum, Scattering | Seattern (@)t | Conter
eyl Rsmvali b s Hodoscope | width (m)
) Degrecs

13000 61018 1405 02806 04m3 0719 00ul
1 T S N () ) S I
3 3083 62713 0959 165566 00726 04574 07300 00104
4031250 6350 | 02919 16738 0268 | 04507 07195 00100
5 31667 6408 0269 164974 02652 0442 0709  0.009
632083 65 0241 16272 02616 04380 06995 00096
7 32500 66103 02603 1608 08l 04319 06%0 00093
83917 6% | 0267 ISES 0247 | 0460 | 06807 0.0091
s 3333 678 02732 IS6S0S 02514 04X 06716 O.088

100330 68685 02697 1427 | 02481 0447 0GRS 00086

1 34167 69493 0266 1297 02450 04090 0652 00083

1 3asE3 0 70M0 02630 10714 | 0240 04039 06459 00081

13 35000 7187 0298 4S8 02390 038 0677 00079

160307 72085 0267 147087 | 02361 0338 | 0698 00071

35 35833 7B 0237 145H9 0232 03889 021 0007

16 3@ 73002507 14362 02305 0341 | 06l46 00073

1 36667 74T 027 14195 0278 035 08T 0002

1837083 7505 0249 140337 | 0251 0319 0600 00070

33 37500 76T 0242 1387 0205 035 0591 00068

037907 7720 024 172 0200 0368 05862 00067

21 3833 7767 02368 13562 017 0360 0SS5 00065

2 3E0 78IS 0342 LAISS | 022 0359 0570 00064

23 35167 798 02317 1327 028 03538 05667 0008

26 35583 BOSI0 0292 | 13130 02105 03499 05604 00061

25 40000 RIS 0267 12918 02083 03461 05543

Total widih of the FSF hodoscape. 018753

Table 1: Detailed summary of the fine spaced forward (FSF) hodoscope. In this
design 41.7 MeV/c momentum bins are used. The radius of curvature in the
magnet, the bend angle in the magnet, the x-displacement at the end of the
magnet, and the x-coordinate at thehodoscope are computed for the momentum
correspondingto the edge of every momentum bin. The counter widths are
computed in theright column.




 Table 1 and 2 summarizes the important parameters for the FSF  arm. The FSF hodoscope is a single layer close packed scintillator hodoscope with fine segmentation covering the momentum range between 3 and 4 GeV/c. It will consist of 24 scintillator counters, each with a momentum acceptance of 1 GeV/24 = 41.7 MeV. A constant momentum acceptance per channel requires adjusting the width of each counter. For every momentum bin the curvature in the magnet was computed and the bending  angle and displacement at the exit of the magnet determined. Using the bending  angle, the horizontal displacement due to the 1.5 m long drift space is computed and  the horizontal position of the particles on the hodoscope found. The hodoscope is 19.75 cm wide and the width of the individual counters range from 6.1 to 11.1 mm. With this design the high momentum counter has a width of 6mm which is roughly twice the photon beam spot size on the converter. The effect of the beam spot size on the FSF  resolution was estimated to be 5 to 13%. A careful MC study of the detector resolution taking into account the spot size, multiple scattering, and the pair production opening angle needs to be performed to optimize the design. Another important factor to be considered when selecting the momentum binning of the FSF is the impact this binning has on the ability to determine the polarization of the photon beam by fitting the shape of the photon energy spectrum. 
[image: image3.jpg]1 3007 L6667 120281 07519 101038
2 o 30833 L6667 120281 07408 107690
3 30833 31250 L6667 D08 07300 104495
4 31250 31667 L6667 12081 07195 101446
s 31667 32083 L6667 nos omss s
5 32088 32500 L6667 12081 [
7 32500 32017 L6667 1081 060 | 9307
s 30017 3333 416667 12081 0687 90
] 333 33750 L6667 08 0676 8507
10 33750 34167 L6667 120281 [
i 34167 34583 L6667 nost o6 83480
2 34583 35000 L6667 120281 [
5 35000 35417 416667 2081 o0& Tom
14 35417 35833 416687 120281 [
15 3583 36250 L6667 1208 0wl 755
16 36250 36667 L6667 120281 06145 73478
i 36667 37083 L6667 081 0sm 7169
18 37088 37500 L6667 120281 06000 6958
19 37500 37917 L6667 no8 0531 6833
n 3717 38333 416667 12081 0se  66m1
2 3E3 38750 L6667 108 055 6505
2 38750 39167 L6667 12081 om0 62
5 39167 358 L6667 120281 03667 6291
% 3958 40000 L6667 12081 05604 60508

Tahle 1: Summary of the FSF hodoscope design. The momentum range, momentum
uncertainty, position, and size for each counter are shown.




 As was evaluated  from Figure 2,  already a 15 bins per GeV gives a fairly accurate determination of the shape( fluctuations in the maximum value due to binning doesn’t  exceed 2-3% with impact less than 1% in polarization value) , while in the present design 24bins/GeV is considered. As part of the MC studies the impact of changing the binning on the uncertainty in the linear polarization measurement should be studied.  

[image: image4.png]x 10°
1000

photans/s/GeV
o
o
o

D
o
o

V\\‘\\\‘\\\‘\\\‘\\\

/
f
/
/
15points/GeV f/ i
p e \ ‘ ‘
[ 3.4mm collimator . L_%/\/«;/\
\/
7‘ | | L L ‘ | | L L ‘ | | | | ‘ L | | | ‘ | | | | ‘ L | | |
6 7 8 9 10 11 1

photon energy (GeV)




            Figure 2 A typical photon energy spectrum is shown for a 3.4mm collimator. Superimposed on the curve are the energy measurement points corresponding to a FSF hodoscope with 15 scintillators.  
The WSF hodoscope is a set of six narrow scintillators positioned to detect electrons in steps of 1GeV/c ranging from 3.25 to 8.25 GeV/c.  The 3.25 GeV/c lower limit for the WSF hodoscope is determined by the width of the magnet poles; a 2.25 GeV/c particle will not pass through the magnet. Signals from the two detector arms are taken in coincidence and the photon energy is the sum of the positron and electron energies, so this combination of detectors then covers the momentum range of 6.25 to 12.21 GeV/c. As the FSF and the WSF measure the positron and electron energies independently, the uncertainty in the photon energy is the sum of the energies of the electron and positron added in quadrature. The expression for the photon energy uncertainty is then:

                           σγ =  (σ +2+  σ-2)1/2
where σγ  is the uncertainty in the photon energy and σ + and σ- are the uncertainties in the positron and electron energies respectively. Correspondingly, the expression for the energy resolution using the relation Eγ = E-+E+ is: 

                        (σγ/ Eγ) =  (σ +2+  σ-2)1/2 /(E-+E+).

It is clear from the above expressions that once the segmentation and therefore the resolution of the FSF has been chosen the optimum design of the WSF is highly constrained.  In this design the widths of the WSF scintillators were  adjusted such that the uncertainty in the photon energy is constant for all combinations of FSF and WSF counter pairs. The width of the WSF counters were chosen so that the uncertainty in the reconstructed photon energy was equal to ½ the momentum bin width in the FSF or 20.8 MeV, σEγ  =  (FST momentum bin width)/2 = 20.8 MeV  The influence of  resolution affecting   factors, mainly of beam spot size and multiple Coulomb scattering , will be considered below.
As the uncertainty in momentum measurement in the FSF counter is 12 MeV, this requires that the uncertainty in the momentum from the WST is 17 MeV.  In Table 3 the most important parameter of the WSF hodoscope are summarized. The widths of the counter needed to measure the momentum with an accuracy of 17 MeV varies from 1.3 cm at 3.25 GeV to 0.2 cm at 8.25 GeV.  By designing the spectrometer so that the photon energy resolution is equal to ½ the FST momentum bin width, adjacent photon energy measurement points are separated by an amount equal to the uncertainty in the measurement. This avoids any strong overlap between neighboring photon energy measurement points.
	CounterWSF
of WSF hodoscope
	 
Momentum   
    (GeV/c)   
GeV/c
	Electron Radius(m) angle

deg
	Radius f  Radius(m)
M
	Distance Displacemto beam linto beamline,

    (cm)
	 Scintill  Width(cm) for σγ=21.7 (17MeV/c)

	WSF1
	3.25
	16.03
	6.62
	68.8
	1.25

	WSF2
	4.25
	12.19
	8.66
	51.9
	0.72

	WSF3
	5.25
	9.84
	10.69
	45.4
	0.51

	WSF4
	6.25
	8,25
	12.75
	35.0
	0.33

	WSF5
	7.25
	7.11
	14.78
	30.1
	0.25

	WSF6
	8.25
	6.24
	16.83
	26.4
	0.19


Table 3: Detailed summary of the wide spaced forward (WFS) hodoscope detectors.

Looking at the widths of the strips in the WSF arm shown in Table 3  column 6, one sees that the width of the counters decreases as the momentum increases. This is necessary to match the momentum spread in the WSF to the photon energy resolution given the fixed resolution in the FSF arm. 
	Eγ range(GeV) 
	    6-7
	   7-8
	     8-9
	   9-10
	  10-11
	11-12

	Widths of WSF strips convoluted with beam spot size.(cm)
	1.29
	0.788
	0.602
	0.46
	0.406
	0.372

	Bin width(MeV/c)
	17.5
	18.6
	20.1
	23.7
	27.6
	33.3

	σγ(corrected)

        (MeV)
	21.8
	22.7
	23.9
	27.3
	30.5
	35.7


Table 4. Summary on WSF strips’ effective width (beam spot  size convoluted) and . 
              impact to a photon energy resolution. The FSF resolution affected by beam 

              spot size is set to 13MeV/c 
Monoenergetic particles produced in the converter do not image to a point on the hodoscope plane due to the several factors, mainly by the beam spot size of 3.4 mm on the radiator, but also multiple scattering in the radiator and the small transverse momentum spread in the pair production process also contribute. The impact of the beam spot size to the photon energy resolution is estimated by computing an “effective” FSF and WSF strips widths size. This effective strip width is computed as a square root of the quadratic sum.. From this strip width the effective momentum uncertainty in the WSF and the photon energy uncertainty are computed. The results are shown in Table 4.  This effect of the position spread nearly doubles the error in the photon energy measurement in the large energy bins. The uncertainty is increased from 21 MeV to 36 MeV for photons between 11 and 12 GeV. In the region of the peak of the coherent bremsstrahlung (P- = 5.25GeV and P+ ranges between 3 and 3.96GeV) the effect of the beam spot size on the energy resolution is 18% (23.9 MeV instead of 21.7MeV).
The details of backward layer hodoscopes are not presented here because the studies to understand if they are needed are still underway and will depend on background rate estimates. However similar trigger hodoscopes are in use in the Hall-B pair spectrometer and are important to suppress the noisy background. Their need in the Hall-D pair spectrometer  may be also be indirectly confirmed from Hall B’s data on single arm and coincidence rates. The geometry of the backward layer hodoscopes is shown in the layout above and is used for the design of the  trigger scheme presented below. The present backward hodoscopes consist of 6 strips behind the WSF hodoscope and 3 strips behind the FSF hodoscopes. The trigger hodoscope arrays are located 1m downstream of the WSF (6) and FSF (24) hodoscopes.
f) Pair spectrometer projected performance
a. Energy resolution, efficiency, and rates

The pair spectrometer energy resolutions for FSF granularities of 16 and 24 channels after correcting for the beam spot-size are summarized in 5. 
	
	Eγ = 8-9 GeV
	Eγ= 11-12GeV

	
	n=16
	n=24
	n=16
	n=24

	σγ(MeV)
	33.1 
	23.9
	43.3
	35.7


Table 5: The uncertainty in the reconstructed photon energy is shown depending on whether the FSF is constructed of n=16 of 24 counters. The uncertainty is shown for both the coherent bremsstrahlung region (8-9 GeV) and the high energy bins. 
There is a noticeable gain in photon resolution below 10GeV for FSF granularity n=24 while at the high energy end of the bremsstrahlung spectrum the beam spot size has a larger effect and so the improvement in resolution is smaller. If one increase  the widths of  WSF counters 5 and 6 to 0.4cm in order to increase the acceptance and therefore count rate, the photon energy resolutions at 12 GeV become 49.6 and 41.7 MeV for n=16 and 24 bins respectively. Detailed studies of the rate as a function of radiator thickness will be performed to determine if the 2 mm counters are acceptable.  

The efficiency for each photon energy bin is defined by a two arm coincidence and has to be accurately determined. The relative efficiency of each pair spectrometer channel can be directly measured using a bremsstrahlung spectrum that has a known shape. An amorphous radiator installed in the goniometer will provide a collimated purely incoherent energy spectrum from which the relative efficiency for each channel can be determined by comparing the measured and calculated energy spectrum. This procedure should be regularly repeated in order to monitor and correct any possible instability. The absolute efficiency of the pair spectrometer can be measured experimentally at a known but small flux using a total absorption counter.
The rates in the pair spectrometer were evaluated as the product of the momentum phase space coverage in both arms at a few photon energies. The acceptance was found to be approximately 0.1% (fraction of registered  in coincidence e+e- pairs for each converted photon)  on average, decreasing toward the high energy end of the bremsstrahlung spectrum (0.15% at 6GeV, 0.1% at 9GeV and 0.08% at 12GeV). With this acceptance and assuming 100% detector efficiency, the expected rate in the region of the coherent bremsstrahlung peak is on the order of 100 Hz for a converter thickness of ~10-3 radiation lengths and a photon flux of ~108 photons/GeV. The higher rate might be  obtained by increase of  converter thickness with cautions against increase of multiple Coulomb scattering contribution in the PS resolution. 
For evaluation below are presented  for comparison:

a) The  relative values of multiple scattering cone spot size (horizontal projection angle's  

             RMS times 3.5m converter distance to FSF and WSF arms ). They are  ranged :

             from  0.2 0 to 0.40mm for thicknesses x=10^-3 to 4x10-3 Xo and Pe=8GeV/c and
             from  0.45 to 0.89mm for thicknesses x=10^-3 to 4x10-3 Xo and Pe=3.5GeV/c

     b) The coordinate resolutions of FSF and WSF arms in 24x6 PS's  configuration::

           are  0.58mm for Pe= 8GeV/c  and 1.7mm for Pe=3,5GeV/c respectively.
 As one may see, the size of multiple scattering  remains still lower than the beam spot size,1mm RMS and  resolutions of PS hodoscopes. This feature might be useful  for upgraded  CEBAF  endpoint energy measurement for PS energy scale calibration using 1-1.5mm wide and 4x10-3 Xo thick converter ,allowing to decrease the beam spot size and perform the measurements  with PS energy resolution in this energy zone app.30-35MeV.. 
  A resolution of  the measured endpoint energy is expected to be app.12-25MeV, that might be comparable  to CEBAF beam energy knowledge accuracy of  10-3 .
The overall error on the photon energy measurement will depend on the of the scintillator dimensions precision and  their positioning accuracy , the precision of the field map, multiple scattering in the converter and the beam spot size on the converter. Each of the contributions to the error will need to be studied in detail.

b) Tagger calibration

 As an instrument with simple configuration the energy of the photons can accurately be determined using the pair spectrometer (PS). Therefore it can be used as a reference to cross calibrate  the tagger spectrometer which has more complicated configuration and operates at a much higher rate (>1011).  The tagger spectrometer measures the energy of the photons in the beam by measuring the energy of the scattered electrons and the pair spectrometer measures the energy of a small fraction of these photons which undergo pair production in the PS radiator. By looking at the events where a coincidence exists between the tagger and the pair spectrometer these two energy measurements can be compared. A typical cross-check of the photon energy measured in the PS and the tagger is shown in a scatter-plot.. Each point in the plot is the combination of PS and tagger measurement. The error bars on the plot correspond to the energy resolutions for the PS and tagger in each energy bin. In this plot the energy resolution of the tagger is higher than that of the PS but this does not affect the precision of the cross-check as enough statistics can be collected to make the uncertainty in the mean of the distributions arbitrarily small.  Random coincidences between the tagger and pair spectrometer could shift the calibration curve so care needs to be taken in the analysis.
c) [image: image1.jpg]Bending Magnet

WSB
Bmax=1§4T

Vacuum Chamber wsF

Left=183em Lve=tsoem  FSF FsB
hodoscopes



Photon beam linear polarization determination
Careful analysis of the shape of the coherent bremsstrahlung spectrum allows one to determine the linear polarization of the photon beam. The methods developed for this task are quite advanced and require precise data over a wide enough energy range above and below the peak region (Insert a reference to your NIM article and the Tubingen code). For example, at if the coherent bremsstrahlung peak is at 9 GeV, the range from 6 to 12 GeV is quite appropriate for this task. The expected precision will be on the order of 0.01-0.02 over the working range of (E-Epeak)/ Epeak<0.3. Outside of this range the polarization is low and compatible with zero.  
[1] Tubingen Group Paper(will find… )
[2] S.Darbinyan, H.Hakobyan,R.Jones, A.Sirunyan and H.Vartapetian , NIM  A554(2005)75-84 
d) Electronics
The entire pair spectrometer consist of a total of 42 counters: FSF(24),FSB(4),WSF(6),and WSB(6). The original proposed readout and trigger schemes for the detector are shown in Figures 4. In Figures 5 and 6 the implementation of this scheme in GlueX standard electronics is presented. The scintillators are read out using photomultiplier tubes which are powered by a CAEN SY2527 based high voltage system. The signals from the FSF and WSF are discriminated using CAEN V812 constant fraction discriminators and [image: image5.jpg]Flash
TDC [—*Pipeline

Flash TDC|

—b|
L [FlashTDC '




Figure 4: Readout scheme for the pair spectrometer.

then readout using the fTDC. The backward layer hodoscopes are used to form the pair spectrometer trigger and are readout with the fADC250 operating in “HIT BIT Mode”. A bit pattern containing the information as to which of the back hodoscope channels was hit is formed in the fADC250 and passed to the crate trigger processor (CTP). The CTP passes this information to the subsystem processor (SSP) and the SSP passes it to the global trigger processor (GTP). The GTP checks for coincidences between the WSB and FSB hodoscopes and generates a pair spectrometer trigger if one is found. 
e) Software                     

A large amount of software will be needed to operate the PS and analyze the data. The major software packages include:

· The software needed to measure the channel by channel efficiency of the PS using the incoherent bremsstrahlung spectra. 

· Software to search for the crystal axes orientation in the radiator using angular scan information. Online software to measure and monitor the coherent bremsstrahlung spectra and provide feedback information for the radiator crystal positioning.

· Off-line pair spectrometer data analysis and coherent bremsstrahlung polarization calculations.
f) Work to do: Next steps

1. Once the magnet is chosen, the precise trajectory calculations need to be performed to determine the optimum detector layout using that magnet field topography. 
2. A full simulation of the pair production process for 6-12 GeV photons using the selected PS layout needs to be performed to determine the precise  layout ,resolutions and efficiencies.

3. The counter structure needs selected: scintillator type, light guides, PMT’s.
4. The PS detector mechanical infrastructure must be designed: hodoscope counter construction including mounts and magnetic shielding, movable supports and rails for precise adjustments of the counters to relevant coordinate references.
5.  Design and prototyping of the counters with special attention to quality control issues. Preparation of the test benches for certification of hodoscopes and counters after assembly. 




Figure � SEQ Figure \* ARABIC �1�: Cross calibration of the tagger using the pair spectrometer.


























12

