## Slide 1: Excimer Laser Progress Report
Brendan Pratt
University of Connecticut
Advisor: Dr. Richard T. Jones

## Slide 2: Recap
Acquired all components required for initial testing.
Initial pulses showed lasing for approximately 5 minutes.
Discovered correct procedure for passivation from a Lambda Physik Engineer.
Duration of light produced from passivation indicates the quality of the laser system.
Laser intensity increased to 140mJ, but duration still lasted only 5 minutes.

## Slide 3: Plan of Attack
After talking with a local laser group, it was decided that the interior of the gas cavity may be contaminated.
We decided to undergo a major disassembly to remove any debris that could contaminate the fluorine gas.

## Slide 4: Inside the Gas Cavity

## Slide 16: Results
We have now seen red light for over 4 hours
The laser beam has not yet been profiled quantitatively.
However, lasing can be visually seen for over 25 minutes.
We plan to continue passivation procedures until we hit a limit.